## Slide 1
Problem: LC898124EP3XCBD# 2 detach component | Machine model: PCP303SL | Machine No: PB011/TS410 | Device: LC898124EP3XC | Package: WLCSP27 | Wafer lot # 5064798010 | Date: 2019.07.24
1) Details of the problem
2) Change Point – Set-Up
3) Hypothesis
4) Containment Actions ( Hypothesis Validations)
LC898124EP3XCBD# 2 detach component
  Missing component at BD#2
Who | Change point | When | Remarks
Mfg. | Running Production | 2019.07.23 | Machine is on running production
Mfg. | Intermittent cont. fail | 2019.07.23 | LSG wait
EE | > Reheat gun probe card #1 > use probe card#2-(DGS error occur) > change socket > brush socket pins and clean probe card pads using bemcot and methanol > realign Test head > reteach APA > series of adjustment on contact height and pad position. | 2019.07.23 | Still encounter intermittent cont. fail
HW | For probe card#1 and 2 checkout | 2019.07.23 | Already complete the missing component- good after completion of parts
Hypothesis | Component | Valid | Invalid | Remarks
Probe Card | Probe Card detach component | x |  | For Completion of parts and component
No | What | Action | Who | When | Status | Remarks
Docking Alignment |  |  |  |  |  | 
Probe card/ load board |  |  |  |  |  | 
1 | Probe card #2 | Install diode and capacitor (S911_D10, S02_D10, S15_D5, and S02_C8) | Renie boy & Dex | 19.07.23 | Closed | Pass after repair
Tester |  |  |  |  |  | 

## Slide 2
7) CAPA and Recommendations:
5) Root Cause Analysis
Problem | Why1 | Why2 | Why3
Detach component at probe card | Possible parts and component hit during set-up, repair or isolation | Mishandling of load board during set-up and repair | 
What | Who | When | Remarks
Need orientation on all group on proper handling of LB, materials and peripherals during repair and set-up. | LSG, Tester, Handler, HW Group | TBD | 
6)Pictures and Illustrations
LC898124EP3XCBD# 2 detach component